## Slide 1: GEO600 Beam Tubes
Harald Lück
Leibniz Universität Hannover
&
Max Planck Institut für Gravitationsphysik
Albert Einstein Institut

## Slide 2: GEO600 vacuum system
Dimensions:
  Volume:400m3/ Surface:4000m2
  Tubes: 2x600m,  60cm
  11 Vacuum Tanks: 2m tall, 1m 
Pumps:
  5 Scroll (25 m3/h, 10-2 mbar)
  4 Turbo-Pumps (maglev, 1600 l/s for N2, 1000 l/s for H2 )
  2 Turbo-Pumps (maglev, 180 l/s)
Pressures:
  Tubes : < 1*10-8 mbar
  Tanks : 3 *10-9 - 3*10-8 mbar

## Slide 3
Vacuum Tanks
Dimensions:
  2m tall, 1m  (9x)
  0,7m (2x)
Materials:
  Stainless 1.4404 ,
  1.4435, 1.4429 (Flansche); glas bead blasted
gaskets:
>250mm COF
 250mm Conflat
24h „Air-bake“ @200°C
„Vacuum-bake“ 48h @ 80°C
Outgassing rates ca. 1E-13 mbar l/s cm2

## Slide 4
Beam Tubes
Stainless Steel 1.4429 (316LN), cold rolled
60 cm ID
2* 133* 4,5m segments (~85kg each) [8E-4*pi*0.6*4.5*8000*pi/2]
0,8 mm wall thickness
30 mm wave period; semi-circular
total weight 2* 12 t
Notes: 316LN 1.4429 (UNS S31653) is a lowcarbon (1.4404, 316L), nitrogen-enhanced version of Type 316 molybdenum-bearing austenitic stainless steel. The Type 316 alloys are more resistant to general corrosion and pitting/crevice corrosion than the conventional chromium-nickel austenitic stainless steels such as Type 304. They also offer higher creep, stress-rupture and tensile strength at elevated temperature. The nitrogen in Type 316LN adds additional resistance to sensitization in some circumstances.

## Slide 5
Making the Beam Tube segments
Manufacturing steps:
Bend to form tube, weld longitudinally
Corrugate
Clean, leak test
Lip weld sections
3cm

## Slide 6
Installation of the beam tubes
Transport (> Mai 1996)

## Slide 7: Joining the segments

## Slide 10: A year of welding

## Slide 11
Installation of the tubes
600m rail
(later used as current return path)
June 1996

## Slide 12
Baking the tube
Isolation: 20cm rockwool Heating: 300V/600A DC 0.5 Ohm/ 600m  800 mΩ/km
= 180kW  300W/m
Passivation in air (dried with silica gel):
2 days @ 200°C
Vacuum bake:
5 days @ 250°C

## Slide 13: Insulating the tube

## Slide 16: Baking
Emergency Diesel 350kW

## Slide 17: Rectifier & Current control

## Slide 18
In situ Leak check
Wrap 3m section of plastic foil around the tube
and flush with He. Observe with RGA @ tube ends. Reaction time < few seconds. Very sensitive…
5 leaks in total.
Fixed with Ceramabond. Some welded? TBC
two subsequent leaks after ca. 20 years:
MIC on top of tube. Marten droppings.
Corrosion @ SS304 (1.4301) hand valve in trench 304 much less corrosion resistant than 316LN under GEO600 trench conditions!

## Slide 19
Pros:
weight 
  simpler handling
  simpler mounting
less material -> low costs / good material affordable
axial elasticity  no need for bellows
high electrical resistance  lower current for baking required for same power/length
GEO600 Beam Tubes
Cons:
not bulletproof (US criterion)
factor p/2 surface increase (can be lowered by shallower corrugation)
More challenging to clean in situ due to corrugation

## Slide 20
Pressure in East Tube
27.01.2000
1E-14
1E-13
1E-12
1E-11
1E-10
1E-09
1E-08
1
11
21
31
41
51
61
71
81
91
Mass [amu]
Pressure [hPa]
Partialdrücke im Ostrohr (unkalibriert)
Lecks
Filamente
Rohr
Wasser
RGA only calibrated to total pressure gauge!
Some remaining leaks in tank cluster sections.
Tubes leak free.

## Slide 21
27.01.2000
1E-18
1E-17
1E-16
1E-15
1E-14
1E-13
1E-12
1
11
21
31
41
51
61
71
81
91
Masse [AME]
Ausgasrate [mbar l/s cm2]
Ausgasraten im Ostrohr
Lecks
Filamente
Rohr
Wasser